#### Review 1 Report

# Design and Development of a Context-Aware TinyML Fall Detection System: Inertial Sensor Integration, Model Evaluation, and Interactive Review Dashboard

# Problem Identification & Relevance

Falls remain one of the leading causes of injury-related morbidity, hospitalisation, and loss of independence among elderly individuals worldwide. Falls are widely recognised as a major public-health concern in ageing populations: a large proportion of adults aged 65 and above experience at least one fall per year, and outcomes are frequently worsened not by the fall itself but by delayed detection and delayed medical response, the so-called "long-lie" problem, where an elderly person remains undetected on the floor for an extended period after falling.

Existing wearable and ambient fall-detection systems (surveyed in Section 2) have made significant progress in detecting the binary event of "a fall occurred," using accelerometer/gyroscope data and machine-learning or deep-learning classifiers. However, a fall alert alone does not tell a caregiver or emergency responder how serious the fall was, or what the wearer's physiological state is immediately afterward. This is a critical, real-world gap: not every fall is medically urgent, and treating every alert as equally urgent leads to alarm fatigue among caregivers, while truly severe falls (e.g., with cardiac stress or breathing difficulty) may not be prioritised appropriately.

Why this problem matters: Accidental falls among the elderly represent a critical public health concern. A reliable system combining low-power inertial sensing with lightweight TinyML models offers an affordable, privacy-preserving solution to detect fall events early. Demonstrating working sensor communication, live 6-axis motion acquisition, robust algorithmic benchmarking, and an interactive review dashboard establishes the essential groundwork for transparent, validated wearable fall detection.

# Literature Survey

A structured review of 30 Q1-indexed journal papers (2013–2026, drawn from venues including Sensors, IEEE JBHI, IEEE TBME, Frontiers in Digital Health, Frontiers in Robotics and AI, Measurement, Biomedical Signal Processing and Control, Computers in Biology and Medicine, Journal of Biomedical Informatics, Journal of Medical Systems, Gait & Posture, Future Internet, and Engineering Applications of Artificial Intelligence) was conducted to establish the current state of the art in wearable and ambient fall-detection technology. The survey spans four broad clusters:

- Foundational wearable architectures and classical ML classification of fall vs. activities of daily living (Papers 1, 2, 9, 21, 22, 28, 30).

- Deep-learning and TinyML approaches for on-device fall classification, including ensemble DNNs, CNNs, transformers, and resource-constrained embedded implementations (Papers 7, 8, 13, 17, 18).

- Context-aware detection strategies that use activity/time/fall-type information to reduce false alarms (Papers 14, 15).

- Alternative modalities and comparison points: vision/depth-camera systems, smartphone audio, ambient/smart-home sensing, and EMG-based approaches (Papers 3, 12, 23, 24, 25), plus fall-risk assessment and gait-based risk scoring (Papers 26, 29), and broad literature reviews mapping the field's overall trends and open challenges (Papers 4, 5, 6, 11, 16, 19, 20, 27).

The full paper-by-paper comparison, including each study's existing approach, sensors/data used, ML method, strengths, limitations, and how the proposed project improves on it, is provided in the Comparison Table (Section 5).

Key gap identified: Across all 30 papers, existing systems either focus purely on high-complexity ML models without embedded constraints, or implement rudimentary threshold heuristics with high false alarm rates. Very few works provide an open, transparent pipeline linking live sensor data acquisition, context-aware TinyML model benchmarking on standardized datasets (SisFall), and an interactive publication review interface for rigorous metric evaluation.

# Proposed Approach & Preliminary Feasibility

#### System Overview

The project architecture integrates an ESP32 microcontroller with a 6-axis Inertial Measurement Unit (MPU-6050 accelerometer + gyroscope) communicating via I²C (GPIO 21/22). Live tri-axial acceleration, angular velocity, and resultant acceleration magnitude are streamed to verify motion sensing. Concurrently, a Python-based TinyML development pipeline trains and benchmarks lightweight deep learning models (TinyFallNet), and a comprehensive HTML/JS review dashboard provides interactive visualization of all evaluation metrics.

#### Processing Pipeline

    - Module 1 — Hardware Sensing & Motion Processing: The ESP32 acquires 6-axis IMU measurements (Ax, Ay, Az in g, Gx, Gy, Gz in deg/s) via I²C, continuously calculating resultant acceleration magnitude to track stationary versus dynamic body movement.

    - Module 2 — Context-Aware TinyML Model Development: IMU signals from benchmark datasets (SisFall) are segmented into 256-sample windows with 88 extracted statistical kinematic features. A compact 1D-CNN architecture (TinyFallNet) is trained with Class-Weighted Focal Loss to overcome severe class imbalance.

    - Module 3 — Multi-Metric Evaluation & Publication UI: Models are benchmarked across 26+ clinical and TinyML metrics (Accuracy, Sensitivity, False Alarm Rate, MCC, INT8 footprint). An interactive, self-contained dashboard integrates all tables, confusion matrices, ROC/PR curves, and threshold calibration sweeps for faculty and journal review.

#### Preliminary Feasibility

Technical feasibility: Basic I²C hardware communication, register configuration, and live 6-axis data acquisition on the ESP32 have been fully established. On the software side, TinyML model training, INT8 quantization profiling (20.48 KB footprint), and metric evaluation have been proven feasible on the SisFall dataset.

Economic feasibility: The prototype utilizes low-cost, off-the-shelf components (ESP32 and MPU-6050), keeping hardware expenditure minimal and accessible for academic and open-source replication.

Preliminary risks: Addressed risks include resolving I²C bus communication failures and handling severe dataset class imbalance. Focus is centered on robust offline model evaluation and interactive dashboard validation.

# Novelty Statement, Research Gap & Domain Justification

#### Novelty Statement

This project establishes a unified framework bridging physical motion sensing and transparent machine learning evaluation: (1) verified hardware-level IMU interfacing and live 6-axis motion acquisition on the ESP32, (2) an optimized, lightweight TinyFallNet model trained with Class-Weighted Focal Loss achieving 95.80% fall sensitivity on SisFall, and (3) an interactive publication review dashboard presenting multi-metric comparisons, ablation studies, and threshold calibration.

#### Research Gap Statement

While prior literature demonstrates binary fall classification, systems frequently lack transparent evaluation across clinical metrics like False Alarm Rate, MCC, and Diagnostic Odds Ratio under severe class imbalance. Furthermore, existing research rarely provides an interactive review tool connecting real-world hardware feasibility with rigorous model benchmarking.

#### Domain Justification

With aging populations worldwide, low-cost wearable sensing and transparent AI evaluation are essential to accelerate practical geriatric monitoring. Demonstrating reliable hardware acquisition alongside verified, reproducible TinyML benchmarks ensures both technical validity and high clinical relevance.

# Comparison Table

## Existing Literature vs. Proposed Project

Based on the 30 Q1 journal papers surveyed, we obtain the following:

| N o. | Existing Approach | Sensors / Data Used | ML Method | Strength | Limitation | How our Model Overcomes It | Link |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | General wearable fall-detection system architecture (hardware + software pipeline) | Tri-axial acceleromet er, gyroscope | Tri-axial acceleromet er, gyroscope | Establishes a practical end-to-end wearable framework | Simple thresholding causes frequent false alarms; no context or severity information | Adds a context-aware TinyML classifier plus severity grading and vitals monitoring on top of this baseline architecture | https://pmc.ncbi.nlm. nih.gov/articles/PMC 4346101/ |
| 2 | Supervised ML classification of fall vs. activities of daily living (ADL) | Acceleromet er (+ gyroscope) | Classical ML (SVM / kNN / Decision Tree) | Shows ML outperforms simple thresholds | Classical ML not optimised for embedded/Tiny ML deployment; no severity or physiological follow-up | Deploys an equivalent classifier as TinyML on-device, then adds severity classification and BPM/SpO2 tracking | https://pmc.ncbi.nlm. nih.gov/articles/PMC 4118339/ |
| 3 | Vision-based fall detection via body-part tracking | Depth camera (Kinect-type) | Geometric body-tracking analysis | High accuracy without a wearable | Fixed installation, privacy concerns, not portable outdoors | Fully wearable and portable, preserves privacy, and adds physiological context the camera cannot capture | https://research.polyu .edu.hk/en/publicatio ns/fall-detection-based-on-body-part-tracking-using-a-depth-camera/ |
| 4 | Review of wearable fall-risk assessment techniques | Various IMUs (review scope) | Mixed (review) | Broad synthesis of risk-assessment approaches | Review only, no working severity-scoring or real-time embedded system | Implements a functioning real-time severity + risk pipeline in hardware, not just literature synthesis | https://www.frontiersi n.org/journals/digital-health/articles/10.338 9/fdgth.2022.921506/ full |
| 5 | Survey of elderly-specific fall detection methods | Various (review) | Mixed (review) | Domain-focused, broad method coverage | No novel system; context-awareness gap identified but unaddressed | Directly implements the context-aware solution the survey calls for | https://www.frontiersi n.org/journals/robotic s-and-ai/articles/10.3389/fro bt.2020.00071/full |
| 6 | Review of wearable fall-recognition systems and open challenges | Various wearable IMUs | Mixed (review) | Clearly frames challenges: power, false positives, context | No implementation offered | Practically resolves the cited challenges via low-power TinyML + context + severity design | https://dl.acm.org/doi/ 10.1016/j.engappai.2 023.105993 |
| 7 | Ensemble deep neural network for fall classification | Wearable acceleromet er / gyroscope | Ensemble DNN | High classification accuracy | Computationall y heavy, not suited to microcontroller deployment; no severity/vitals output | Uses a compressed, MCU-deployable TinyML model and extends output to severity + vitals | https://www.ncbi.nlm. nih.gov/pmc/articles/ PMC10223485/ |
| 8 | TinyML CNN-based wearable fall detector | IMU (acceleromet er/gyroscope ) | Tiny CNN (embedded, quantised) | Proves feasibility of on-device TinyML fall detection | Binary fall / no-fall output only; no severity grading or vitals | Extends the same TinyML approach with severity classification and post-fall BPM/SpO2 monitoring | https://www.scienced irect.com/science/arti cle/abs/pii/S1746809 423007589 |
| 9 | Feature-engineered accelerometer data classified via ML | Acceleromet er (single modality) | Classical ML (Random Forest / SVM) | Good baseline accuracy with interpretable handcrafted features | Single-sensor input; no gyroscope/cont ext fusion; no severity output | Fuses accelerometer + gyroscope + context, then adds severity and vitals layers | https://www.mdpi.co m/1424-8220/20/22/6479 |

| 1 0 | IoT-enabled fall detection with location-based alerting | IMU + GPS/location module | On-device / cloud classifier | Adds actionable emergency response (location + alert) | No physiological severity assessment of the detected event | Adds BPM/SpO2-based severity triage alongside the location alert for better-informed response | https://www.mdpi.co m/1424-8220/25/12/3632 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 1 | Bibliometric/ne twork review of a decade of wearable fall-detection research | Review (various) | Review (various) | Maps 10-year research trends and persistent gaps | No new system; confirms the integration gap between detection, severity and vitals | Directly targets the exact integration gap this review identifies | https://www.mdpi.co m/1424-8220/25/7/2205 |
| 1 2 | Ambient-assisted-living and smart-home based fall detection | Ambient sensors, cameras, smart-home IoT | Mixed | Effective for passive indoor monitoring | Not portable outside the home; no wearable physiological tracking | Fully wearable and mobile, functional indoors and outdoors, with onboard vitals monitoring | https://www.mdpi.co m/1424-8220/25/21/6540 |
| 1 3 | Real-time fall detection optimised for resource-constrained devices | IMU | Lightweight / TinyML classifier | Demonstrates feasibility on constrained MCUs | No severity grading or health-vitals extension | Builds on this constrained-hardware feasibility and adds severity + BPM/SpO2 monitoring | https://www.mdpi.co m/1999-5903/18/4/211 |
| 1 4 | Differentiates hard (rapid) vs. slow falls using context | IMU | Context-aware classification algorithm | Reduces false alarms by distinguishing fall types | No physiological/s everity follow-up after classifying fall type | Combines fall-type context with a full severity score and post-fall vitals monitoring | https://ieeexplore.iee e.org/abstract/docum ent/10592536 |
| 1 5 | Uses time-of-day/activity context with inertial data to reduce false alarms | IMU | Rule-based context-aware algorithm | Early proof that context reduces false positives | Limited context types; no severity or vitals dimension | Uses a richer, TinyML-based context model plus severity and vitals monitoring | https://pubmed.ncbi.n lm.nih.gov/28268395/ |
| 1 6 | Review of wireless communicatio n + AI approaches for elderly fall detection | Review (various, IoT) | Review (various AI) | Covers communicatio n and AI integration trends | Identifies but does not solve the severity-assessment gap | Implements the missing severity + vitals layer on top of standard IoT alerting | https://www.scienced irect.com/science/arti cle/abs/pii/S0263224 124000708 |
| 1 7 | Transformer-based deep learning on IMU data patches | IMU | Patch-Transformer network | State-of-the-art accuracy via attention mechanism | High compute/memo ry cost, unsuitable for TinyML MCUs; binary output only | Uses a TinyML-compatible lightweight model and outputs severity + vitals, not just a binary label | https://www.mdpi.co m/1424-8220/23/14/6360 |
| 1 8 | Embedded on-device fall detection robust to sensor placement | IMU (position-independent) | Embedded classifier | Robust across different on- body sensor placements | No severity or vitals extension | Retains placement flexibility while adding severity classification and vitals monitoring | https://www.mdpi.co m/1424-8220/19/11/2426 |

| 1 9 | Review comparing sensor types/placeme nts for fall monitoring | Review (acceleromet er/gyroscope at various placements) | Review | Guides optimal sensor choice and placement | No implementation ; no severity/vitals discussion | Applies these placement findings directly when designing the proposed wearable | https://www.mdpi.co m/1424-8220/18/5/1613 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 2 0 | Survey of algorithms, datasets and challenges in wearable fall detection | Review (various) | Review (various) | Broad coverage of algorithmic advances | No new system; confirms the severity/vitals gap | Fills this confirmed gap with a working, deployed solution | https://pmc.ncbi.nlm. nih.gov/articles/PMC 7201510/ |
| 2 1 | Personalised single-sensor IoT system for fall prediction and detection | Single IMU sensor + IoT connectivity | Patient-specific model | Personalisatio n improves per-user accuracy | Single sensor limits context richness; no severity/vitals output | Adds multi-modal context sensing and severity/vitals output over a similar IoT backbone | https://pmc.ncbi.nlm. nih.gov/articles/PMC 7201510/ |
| 2 2 | ML validated on real-world (not simulated) elderly fall data | Lumbar-worn acceleromet er | ML classifiers on real fall dataset | Real-world validation increases clinical relevance | No severity grading; single placement; no vitals | Builds on real-world-validated detection and adds severity grading plus vitals monitoring | https://pubmed.ncbi.n lm.nih.gov/28269098/ |
| 2 3 | Combines EMG and accelerometer signals for activity/fall monitoring | sEMG + acceleromet er | Signal-processing/c lassification framework | Multi-modal sensing improves robustness | EMG adds cost, complexity and discomfort for elderly users; no severity/vitals | Uses a simpler, comfortable IMU + PPG combination to deliver severity and vitals instead | https://pubmed.ncbi.n lm.nih.gov/24234563/ |
| 2 4 | Detects falls via smartphone microphone/au dio features | Smartphone microphone | Audio feature classifier | Novel modality, no wearable required | Sensitive to ambient noise; needs phone proximity; no vitals output | Uses reliable wearable IMU-based detection with physiological monitoring, independent of ambient sound | https://pubmed.ncbi.n lm.nih.gov/25915965/ |
| 2 5 | Vision-based deep learning (3D-CNN + LSTM) on video kinematics | Camera / video | 3D-CNN + LSTM | High accuracy via spatiotempor al modelling | Requires fixed camera field of view, raises privacy concerns, heavy compute | Privacy-preserving, camera-free wearable TinyML alternative with added vitals monitoring | https://pubmed.ncbi.n lm.nih.gov/29994460/ |
| 2 6 | Deep-learning gait analysis for fall-risk classification | Wearable IMU (gait) | Deep-learning gait classifier | Introduces fall-risk (not just fall-event) scoring | Focused on a specific clinical population/gait; no post-fall severity or vitals | Extends the risk-scoring concept to elderly users and adds post-fall severity + vitals monitoring | https://pmc.ncbi.nlm. nih.gov/articles/PMC 8221405/ |
| 2 7 | Systematic review of fall detection and fall-risk assessment evidence | Review (various) | Review | Consolidates the wearable risk- assessment evidence base | No implementation ; no post-fall severity assessment | Implements the evidence-based design as a working severity-assessment system | https://www.scienced irect.com/science/article/abs/pii/S1386505 619303375 |

| 2 8 | Multi-sensor fusion via covariance-matrix features for classification | Multiple wearable IMUs | Covariance-matrix-based classifier | Robust fusion improves detection accuracy | Multiple sensors increase cost/discomfort ; no severity/vitals output | Achieves comparable robustness from a single compact wearable, plus severity and vitals | https://www.scienced irect.com/science/arti cle/pii/S1532046419 301078 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 2 9 | Smartphone-based postural-sway measurement for fall-risk screening | Smartphone IMU | Statistical postural-stability metrics | Low-cost, accessible risk screening | Screening only : not real-time fall-event or severity detection | Combines the risk-screening concept with real-time detection, severity grading and vitals | https://www.scienced irect.com/science/arti cle/abs/pii/S0966636 218304521 |
| 3 0 | Hybrid threshold-analysis + ML algorithm for real-time detection | IMU | Hybrid rule-based + ML classifier | Balances computational efficiency with ML accuracy | No severity classification or vitals monitoring after detection | Adds a severity-assessment stage and BPM/SpO2 monitoring after the hybrid detection step | https://pubmed.ncbi.n lm.nih.gov/36085845/ |

#### Existing Product vs. Proposed Project

Based on the existing products, we analyse and observe the following

| Vendor – Product | Model / Algorithm Used | Research Gap / Limitation | How Our Proposed Model Overcomes It |
| --- | --- | --- | --- |
| Apple – Apple Watch Fall Detection | Proprietary on-device hard-fall detection using motion sensors; detects hard-impact events and checks user response/immobility before emergency calling | Mainly designed for hard-fall detection and SOS workflow; internal ML model and decision features are not transparent; alert mainly provides emergency response/location rather than an explainable severity score with post-fall vital signs | Uses a transparent, deployable TinyML model on ESP32-S3/STM32; classifies activity before confirming a fall; calculates a sensor-derived emergency-priority level; sends BPM, SpO₂, location/context and fall confidence to caregiver |
| Samsung – Galaxy Watch Hard Fall Detection | Proprietary hard-fall detection based on watch motion data; SOS and emergency-contact notification workflow | Primarily concentrates on hard-fall events; limited publicly described context-aware analysis; no open model for testing/tuning; no dedicated caregiver output combining severity, prolonged inactivity and physiological trend | Adds multi-stage context awareness: activity type, time, room/location, orientation change and inactivity period; combines this with BPM and SpO₂ trends to send a detailed caregiver alert |
| Lively – Lively Mobile2 | Patented/proprietary accelerometer algorithm; analyses sudden movement relative to physical activity and posture | Device positioning can affect detection performance; typically focuses on automatic alert/response service rather than local explainable ML, fall-type classification, or physiological post-fall assessment | Uses IMU-based TinyML and can be trained/evaluated for multiple activities and simulated fall directions; adds posture confirmation, fall confidence, context rules, BPM and SpO₂ monitoring; all |

| Vendor – Product | Model / Algorithm Used | Research Gap / Limitation | How Our Proposed Model Overcomes It |
| --- | --- | --- | --- |
|  |  |  | decision logic is visible for academic evaluation |
| MobileHelp – Solo / Duo / Elite / Fall Button | Proprietary automatic fall-detection technology with cellular emergency-monitoring workflow; specific algorithm details are not public | Often service/subscription dependent; primary outcome is emergency call/monitoring; algorithm cannot be customized for a local population, activity type, or academic experiment; limited transparent severity classification | Uses a low-cost open prototype with user-tunable model/thresholds; performs inference locally, sends only alert events, and logs confidence, severity level, activity context, BPM and SpO₂ for evaluation |
| Medical Guardian – Active Guardian / Mobile Guardian | Proprietary fall-detection technology; uses motion/height-change pattern recognition and operator verification | Fall detection may not detect all falls; device may require appropriate wear position; commercial system is designed for dispatch rather than transparent false-alarm analysis and custom TinyML research | Uses activity recognition plus impact, orientation and post-fall inactivity confirmation to reduce false alarms; reports precision, recall, F1-score and false-alarm rate; supports configurable wearable placement experiments and context rules |
| Bay Alarm Medical – SOS Mobile / SOS Smartwatch | Proprietary automatic fall-detection technology integrated with GPS, caregiver tracking and monitoring station | Focus is rapid connection to a monitoring station; limited public information about model behavior, severity reasoning, sensor-fusion method, or on-device intelligence; functionality may depend on subscription/connectivity | Introduces transparent on-device TinyML, sensor fusion, severity reasoning and local alert generation; can still send alerts using Wi-Fi/BLE/GSM but does not continuously depend on a proprietary monitoring service |

# Presentation & Team Coordination

#### Team Roles

All members will jointly participate in system integration, testing, literature review, report writing, presentation preparation, and final project demonstration.

| Team Member Name | Role / Responsibility | Key Deliverable(s) |
| --- | --- | --- |
| Bharath Palanisamy Mettukadai | Hardware Implementation | ESP32 board setup and programming, I²C bus scanning (address 0x68), MPU-6050 IMU interfacing, raw register reading, live 6-axis data acquisition (Ax, Ay, Az, Gx, Gy, Gz), and acceleration magnitude computation. |
| Kaleeshwar | ML Model | TinyFallNet architecture design (Multi-Scale Depthwise SE-ResCNN), Class-Weighted Focal Loss formulation, Python model training & validation on SisFall dataset, loss ablation studies, and INT8 quantization profiling. |
| Hiba Fathima | Embedded Coding & ML Testing / Publication UI | Firmware test sketches and sensor processing routines, 5-model benchmarking pipeline across 26+ clinical metrics, decision threshold sweep, and development of the interactive publication review dashboard (dashboard.html) and README documentation. |

#### Project Timeline

| Phase | Milestone | Target Weeks |
| --- | --- | --- |
| Phase 1 | Literature survey finalized; scope, problem, and methodology locked | Week 1-2 |
| Phase 2 | ESP32 setup, I²C bus scanning, MPU-6050 sensor interfacing, and live motion data streaming | Week 3-4 |
| Phase 3 | TinyML model training (TinyFallNet) on SisFall dataset and INT8 quantization profiling | Week 5-6 |
| Phase 4 | Comprehensive model benchmarking across 5 architectures and loss ablation analysis | Week 7 |
| Phase 5 | Decision threshold calibration sweep and metric parameterization | Week 8-9 |
| Phase 6 | Interactive publication dashboard (dashboard.html) development, documentation, and final review | Week 10-12 |

# Proposed Methodology / Detailed Description of the Work Carried Out (Innovation / Novelty)

### Overview of the Proposed Work

The proposed work establishes an integrated framework for elderly fall detection focusing on verified inertial sensing, rigorous TinyML model evaluation, and an interactive publication review interface. The methodology combines physical ESP32 IMU acquisition, 88-feature statistical kinematic extraction, deep convolutional attention modeling, and multi-metric benchmarking.

### System Architecture

The system architecture integrates the following core hardware and software components:

- 6-axis Inertial Measurement Unit (MPU-6050) – accelerometer and gyroscope – for capturing tri-axial motion, angular velocity, and resultant acceleration magnitude.

- Low-power ESP32 Microcontroller – programmed via Arduino IDE to manage I²C communication (GPIO 21 SDA, GPIO 22 SCL), register configuration, and live sensor data acquisition.

- Python TinyML Development Pipeline – for dataset processing, 88 kinematic feature extraction, model training, and INT8 quantization profiling.

- Interactive Publication Review Dashboard (dashboard.html) – self-contained web interface embedding Chart.js visualizations, benchmark tables, and publication-ready figures for faculty review.

These components link real-world sensor acquisition with rigorous, transparent machine learning evaluation, providing a complete framework for wearable fall detection research.

### Methodology / Processing Pipeline

The core methodology is implemented as a four-stage on-device processing pipeline, summarised in Table 1 and described in detail thereafter.

| Stage | Function | Key Techniques / Inputs | Output |
| --- | --- | --- | --- |
| Stage 1 | Inertial Sensing & Magnitude Calculation | Live 6-axis IMU polling via I²C (50 Hz); raw register conversion to g and deg/s; resultant acceleration magnitude calculation | Live sensor telemetry (Ax, Ay, Az, Gx, Gy, Gz, Mag) |
| Stage 2 | TinyML Fall Detection Modeling | 256-sample sliding windows; 88 statistical kinematic features; TinyFallNet (SE-ResCNN) with Class-Weighted Focal Loss (γ=2.0) | Multi-class classification (ADL, Pre-Fall, Fall) |

| Stage 3 | Multi-Metric Parameterization & Ablation | 5-model comparative benchmark; 6-configuration loss ablation; 19-point decision threshold calibration sweep (τ=0.05 to 0.95) | Clinical metric tables & 300 DPI publication plots |
| --- | --- | --- | --- |
| Stage 4 | Interactive Publication Review UI | Self-contained dashboard.html utilizing Chart.js; embeds benchmark tables, confusion matrices, ROC/PR curves, and Pareto frontier | Interactive faculty review & evaluation dashboard |
| System | Integrated Review Pipeline | Seamless pipeline bridging physical ESP32 motion sensing with comprehensive offline model evaluation | Complete academic & publication evaluation |

Table 1: Stage-wise processing pipeline of the proposed system.

### Stage 1 – Inertial Sensing & Magnitude Calculation

Raw accelerometer and gyroscope signals are acquired from the MPU-6050 via I²C. Accelerometer readings are converted to g, and angular rates to degrees per second. The resultant acceleration magnitude |a| = sqrt(Ax² + Ay² + Az²) is continuously computed to distinguish stationary states (~1.0 g) from dynamic body movements.

### Stage 2 – Context-Aware TinyML Fall Detection

Fixed-length windows (256 samples x 6 channels) from the SisFall dataset are processed to extract 88 kinematic features. The proposed TinyFallNet architecture incorporates multi-scale depthwise convolutions (k=3, 7, 15) and Squeeze-and-Excitation channel attention to reliably separate falls from everyday activities of daily living.

  - Stage 3 – Multi-Metric Parameterization & Ablation Study

To address severe class imbalance (~96% ADL vs ~3% Falls), models are trained using Class-Weighted Focal Loss (γ=2.0). A systematic ablation study compares 6 loss formulations, and a decision threshold sweep (τ=0.05 to 0.95) identifies optimal clinical operating points, balancing fall sensitivity against false alarm rates.

to the raw PPG waveform to obtain reliable physiological estimates even in the presence of residual post-fall motion.

### Stage 4 – Interactive Publication Review Dashboard

All evaluation outputs, confusion matrices, ROC/PR curves, and TinyML hardware profiling metrics (20,975 parameters, 20.48 KB INT8 footprint) are integrated into a self-contained HTML/JS dashboard (dashboard.html). This provides faculty reviewers with an interactive, verifiable interface to examine model trade-offs without requiring complex environment setup.

### Innovation and Novelty of the Work

The innovation of this project lies in connecting real-world hardware sensor interfacing with an exhaustive, publication-grade machine learning evaluation framework:

- Live Hardware Inertial Acquisition – Established verified I²C communication, raw register parsing, and acceleration magnitude calculation on an ESP32 platform.

- Lightweight Depthwise SE-ResCNN (TinyFallNet) – Designed an ultra-compact neural architecture requiring only 20,975 parameters and 20.48 KB INT8 memory.

- Class-Weighted Focal Loss Formulation – Successfully countered 96:1:3 class imbalance, elevating fall sensitivity from 91.41% to 95.80% (and 97.81% under calibrated threshold).

- Interactive Review & Parameterization Dashboard – Delivered a self-contained publication UI enabling comprehensive inspection of all clinical and hardware metrics.

Together, these contributions deliver a verified hardware-to-software pipeline supporting reproducible academic research in TinyML fall detection.

### Implementation Plan

The work was executed in the following structured sequence: (a) Bharath: ESP32 hardware setup, I²C bus scanning, MPU-6050 interfacing, and live 6-axis data streaming with magnitude calculation; (b) Kaleeshwar: SisFall dataset preparation, TinyFallNet model architecture design, Class-Weighted Focal Loss training, and INT8 quantization profiling;

(c) Hiba: Firmware test sketch development, 5-model benchmark pipeline execution across 26+ metrics, loss ablation, threshold sweep, and creation of the interactive publication review dashboard (dashboard.html) and comprehensive README documentation.

### Feasibility Considerations

Technical feasibility is confirmed by verified live sensor data acquisition on the ESP32 and successful completion of the Python training and benchmarking pipeline. Economic feasibility is demonstrated by utilizing readily available, low-cost components (ESP32 and MPU-6050). The integration of an interactive dashboard ensures transparent, verifiable project review.

### Feasibility Analysis

Building on the preliminary feasibility discussion in Section 3.3 and the summary in Section 6, this section presents a more detailed feasibility analysis of the proposed wearable system across technical, economic, operational, schedule, and legal/regulatory dimensions, together with a consolidated view of risk and mitigation.

Technical Feasibility: The primary technical pillars of this project — 6-axis IMU sensing on an ESP32 microcontroller, compact TinyML model development (TinyFallNet), and interactive web-based evaluation — have been empirically validated. Live I²C register reading confirms reliable motion acquisition, while Python benchmarking verifies that TinyFallNet achieves 95.80% fall sensitivity with an INT8 footprint of only 20.48 KB, well within microcontroller SRAM budgets.

Economic Feasibility: All hardware components (ESP32 development board and MPU-6050 sensor module) are low-cost, off-the-shelf parts, costing under $10 combined. Software development and dashboard evaluation utilize open-source frameworks (PyTorch, scikit-learn, Chart.js), requiring zero recurring subscription or proprietary licensing fees.

Social Feasibility: The proposed wearable exhibits high social feasibility by directly addressing key human and behavioral factors in elderly care. Traditional camera-based or ambient monitoring solutions often face strong user resistance due to privacy intrusion. By executing all machine-learning inference locally on-device and eliminating video/audio recording, the system fully preserves personal privacy. Furthermore, by incorporating sensor-derived severity grading alongside physiological monitoring, the device actively mitigates "alarm fatigue" for family members and caregivers by prioritizing urgent emergencies over low-impact events. The lightweight, wrist- or chest-worn form factor ensures high daily wearability and non-intrusive integration into the user's daily living routine.

Environmental Feasibility: The prototype relies on low-power embedded components and standard USB power, minimizing electronic waste and energy consumption during testing. Software evaluation is conducted locally with lightweight compute requirements.

Political / Demographic Feasibility The project is strongly aligned with global demographic shifts and current healthcare policy frameworks. With rapidly aging populations worldwide and a growing ratio of elderly individuals living independently, there is urgent demand from healthcare authorities and municipal bodies for scalable, home-based monitoring systems. From a legal and regulatory perspective, the device's privacy-by-design architecture (on-device TinyML execution and minimal data transmission over BLE) aligns with international data protection mandates such as GDPR, HIPAA, and national digital personal data protection acts. By providing low-cost, decentralized fall triage, the system supports public health policies aimed at reducing emergency hospital admissions and lowering the overall burden on national healthcare infrastructures.

Risk Assessment and Mitigation: Principal risks addressed include hardware communication stability and dataset class imbalance. I²C initialization challenges were resolved by switching to raw register reads, and severe class imbalance was mitigated using Class-Weighted Focal Loss. Shifting focus to a rigorous offline benchmarking and interactive UI framework guarantees high academic integrity and transparent evaluation.